## Slide 1: Metrology

## Slide 2: Metrology
  Instrumentation and Its
  Limits
  Science of physical measurement applied to
  variables such as dimensions
Metrology

## Slide 3: Outline
Measurement using balances
Precision versus Accuracy
Calibration of balances
Validation of pipetmen

## Slide 4: Types of Balances
Labs are equipped with two types of balances:
Analytical balance
Top loading balance
The primary difference between these instruments is Significant Figures

## Slide 5: Precision = +/- 0.01g

## Slide 6: Top Loading Balance
Used when less quantitative results are required (+/- 0.01g) (Capacity < 1200 g)‏

## Slide 7: Analytical Balance
Used for measurements requiring highly quantitative results i.e. +/- 0.0002 g
(Capacity < 100 g)‏

## Slide 8: Bio-Rad Pipetmen
Pipets can measure +/-1% of their largest volume and be accurate

## Slide 9: Limits of Measurement
All measurements contain some error and it is calculated by the following formula:
Percent Error =
True Value - Average Measured Value X 100%
True Value
All

## Slide 10
Standard Deviation Gives a Measure of Variability
 (x - x)2
n - 1
S =

## Slide 11: What are Significant Figures?
The necessary number of figures (digits) required to express the result of a measurement or calculation so that only the last digit in the number is in doubt.
Measuring gives significance (or meaning) to each digit in the number produced.

## Slide 12: Why Consider Significant Figures?
Science depends upon experimentation which requires numerical measurements.
Measurements are taken from instruments made by other human beings.
NO measurement is exact
Error is always a factor

## Slide 13: Determining Significant Figures
Last figure is estimated in measuring 2.33
All whole #'s are significant 2.33 = 3 sig figs
All zeros between 2 numbers are significant 2.03 = 3 sig figs
Zeros to the right of whole number digits are significant if prec by a decimal point
203.00 = 5 sig figs
All zeros to the right of a decimal point & whole number are significant 2.0230 = 5 sig figs
Zeros to the right of a decimal but to the left of a whole number are not significant 0.0203 3 sig figs

## Slide 14: Accuracy
The accuracy of an analytical measurement is how close a result comes to the true value. The analytical method is calibrated using a known standard to determine the accuracy of a measurement.

## Slide 15: Precision
- The reproducibility of multiple
measurements.
- It is evaluated statistically using standard
deviation, standard error, or confidence
interval.

## Slide 16: Measurement Requires Accuracy and Precision
Average value of X

## Slide 17: Accuracy vs Precision

## Slide 18: Calibration of Balances
Balance is reset to detect a specific weight according to directions
A 200.00 g standard is placed on the balance
After recalibrating, the balance will then show a value equal to the standard

## Slide 19: Validation of Pipetmen
Pipetman is selected and set to a specific volume : 1000 μl or 200 μl
Water is drawn up to a desired volume and place into weighed small beaker
Using the equation:
Density= Mass/Volume
*calculate the volume using 1 gram/ml for density of water

## Slide 20: Percent Error
The deviation from an expected value can be expressed as a Percent Error
To calculate the Percent Error use the formula below:
% Error = True Value – Average value x 100
True Value

## Slide 21: Electromagnetic Spectrum
The spectrum of electromagnetic waves ranges from low-frequency radio waves to high-frequency gamma rays. Only a small portion of the spectrum, representing wavelengths of roughly 400–700 nanometers, is visible to the human eye. © Merriam-Webster Inc.

## Slide 22: Spectroscopy
When light of a specific wavelength interacts with a substance, the subsequent energy transfer results in:
    Absorption
    Fluorescence

## Slide 23: Absorption
As the light hits a substance,
energy is transferred to the substance
thereby raising its energy to an excited
state.

## Slide 24: Fluorescence
Molecular absorption of a photon triggers the emission of another photon with a longer wavelength when the molecule relaxes back to its ground state.
Excitation: S0 + hv → S1
Fluorescence: S1 → S0 + hv
h = Planck’s constant
v = frequency of light
S0 = ground state of the fluorescent molecule
S1 = excited state

## Slide 25: Examples of Fluorescence
Molecules that are excited through light absorption can transfer energy to a second molecule, which is converted to its excited state and can then fluoresce.
Fluorescent lights
Light emitting diodes (LEDs)‏
Mercury vapor streetlight
Glow sticks
Compact fluorescent lighting (CFL)‏

## Slide 26: UV/Vis Spectrophotometer
Principle:
Absorption of light in the visible and ultraviolet spectrum results in changes electronic structure of molecules

## Slide 27: UV/Vis Spectrophotometer
Dual light source:
Visible range: Tungsten lamp (400 – 700 nm)‏
UV range: Deutrium lamp (200 – 400 nm)‏
Sample cells
Detector
Mirrors
Grating Monochrometer

## Slide 28: Application
Lambert-Beer Law A = εcl where:
A = absorbance
ε = molar extinction coefficient (L mmol-1 cm-1)‏
c = molar concentration (mM)‏
l = pathlength (cm)‏
The Lambert-Beer law is used to accurately determine the concentration of a substance by measure absorbance at a specific wavelength

## Slide 29: Determining Concentration
The Lambert-Beer Law is used to determine the concentration of an unknown using a standard curve.
y = 0.063 x + 0.002 R2 = 0.998
Equation for a line

## Slide 30: Determining Nucleic Acid Concentration
Quantitative measurements (µg) for nucleic acids (DNA and RNA) at A260
A = εcl
  ε is specific for each type of nucleic acid
  1 OD260 of ds-DNA = 50 µg/mL
  1 OD260 of ss-DNA = 37 µg/mL
  1 OD260 of ss-RNA = 40 µg/mL

## Slide 31: Other Wavelengths
A280: Protein
A230: Phenol and peptide
A260/280: Purity of nucleic acid preparation
(Pure range 1.8 – 2.0)‏

## Slide 32: Terminology
Solution: A homogenous liquid mixture composed of two or more substances
Solute: Substance which dissolves in the solvent
(ex H20)‏
Solvent: Liquid in which solute is dissolved
Concentration: The ratio of the mass of a solute to the volume of solution Examples: g/mL, mol/L, %
Dilution (factor, medium, volume):
Serial dilution is logarithmic
Linear dilution using C1V1 – C2 V2

## Slide 33: Terminology (cont.)‏
Aliquot: Equally divided portions of a sample
Buffer: A salt solution which resists change in pH upon addition of acid or base
Reagent: A substance which is involved in or consumed during a chemical reaction or to detect other substances
Meniscus: A curve in the surface of a liquid which results from interaction with the container

## Slide 34: Common Buffers
Optimal pH ranges for common laboratory buffers

## Slide 35: pH
pH = -log [H+]
pH 7: Neutral: [H+] = 10-7
pH > 7: Basic: [H+] > 10-7
pH < 7: Acidic: [H+] < 10-7
A pH meter is an instrument used to measure the pH of a liquid. Components include:
Probe (glass electrode)
Meter (measures and displays pH)‏

## Slide 36: How pH is Measured
When two solutions with different pH values exist inside and outside a glass membrane, an electromotive force is proportional to the difference between the two pH values. The solution inside the glass membrane has a pH value of 7. The pH value of the solution outside the membrane can be obtained by measuring the electromotive force generated in the membrane.

## Slide 37: pH Meter
The pH meter consists of a glass electrode and a reference electrode. It allows the pH value of the sample to be obtained by measuring the potential difference between the two electrodes with a potential difference meter.
To calibrate the pH meter, a standard solution with a known pH value is used. As standard solutions, phthalic acid (pH 4.01), neutral phosphate (pH 6.86), and borate (pH 9.18) are mainly used.

## Slide 38: Preparing solutions
Concentration is the ratio of the mass (or volume) of a solute to the mass (or volume) of the solution (or solvent)‏
  Mass/volume (%w/v)‏
  Volume/volume (%v/v)‏
  Molar volume (mol/L or M)‏

## Slide 39: Sample Calculation
Ex. Prepare 1 L of 1.5 M Tris pH 7.5
Determine total number of moles of Tris required. 1 L X 1.5 mol/L = 1.5 mol
Tris (MW = 121.1 g/mol)‏
1.5 mol X 121.1 g/mol = 181.65 g Tris
3. Dissolve 181.65 g Tris in 700 mL dH20 (using beaker)‏
4. Adjust pH to 7.5
5. Transfer to 1 L graduated cylinder. Bring to final volume with dH20.

## Slide 40: Material Safety Data Sheet (MSDS)‏
Detailed information on
  Physical and chemical hazards
  Handling procedures
  Emergency response procedures
There must be a MSDS for every chemical used and stored in a laboratory
MSDS for all chemicals must be read and understood before starting a procedure

## Slide 41: Labels
Your first line of defense are container labels
Chemical manufacturers post physical and health hazards on container labels
What information goes on a container label??
  Name
  Manufacturer’s information
  Physical hazards
  Storing, handling, and disposal
  Health hazards
  PPE

## Slide 42: National Fire Protection Association - NFPA CODE
Blue = Health
Red = Flammable
Yellow = Reactive
Scale = 0 – 4
Least Intense to Most Intense
White = Special Hazard
Ex : ACID, NO WATER
http://www.orcbs.msu.edu/chemical/nfpa/nfpa.html

## Slide 43: LABELING… Biohazards
What is a Biohazardous material?
  Biological in nature
  Capable of producing harmful effects
  on other biological organisms, particularly humans
  Examples:
  Certain bacteria, fungi, viruses, parasites, recombinant products, allergens, cultured human or animal cells and their potentially infectious agents